## Slide 1: Light Detection with Ultra-High Dynamic Range
Group 35
DoHyun Kim, Leran Firer, and Eric Kleinberg
Client: Prof. Jon Silva

## Slide 2: Project Review - Client Request
Dr. Silva’s studies conformational and functional change of protein/protein channels by using a fluorescence microscopy.
Notes: Explain/review the fluorescence microscopy again.
1. Components – light input, cell submerged into patch clamp, light through cell and lens collected in PIN diode as the number of photons. PIN diode excited and elicit current. By frequency analysis of the current, models of protein kinase can be established.

## Slide 3: Project Review - Client Request
Background Signal ≈ nA
Signal of interest ≈ pA → SNR is approximately 0.1% → Client requested to build a system with a higher SNR by reducing noise
Reducing Noise
  By matching Lens and PIN diode
  By lowering temperature of the system
  By Input constraint and Filtering system with External circuit

## Slide 4: Project Review: specific design requirements
System | Requirements
Electrical System | PIN Diode Responsivity at 600nm light beam Low Noise Wide Operational Temperature range Minimized Active Area Bandpass Filter of 5 kHz to 10 kHz Amplifier with a gain of 100 or larger Attach the system next to the SM1 adaptor
Cooling System | Peltier or other cooling method to cool PIN diode to optimal temperature (typically around -20 degrees Celsius). Cooling system must fit client’s Sm1 Adaptor.
Light Focusing System | Achromatic doublet lens Known diameter = 25mm Minimized focal length Focus to diffraction limited spot Minimize optical aberrations
Cost | 1000$
Notes: Explain the noise source and what can we do… copy the specific design requiremets once the paper is done.

## Slide 5: Light Focusing System cont’d
Chosen solution: Achromatic Doublet Lens
Notes: completely eliminate the spherical and longitudinal aberration when focusing light to a diffraction limit.

## Slide 6: Lens Specifications
Product Used in Design | Achromatic Doublet Lens |  | 
Manufacturer | Thor Labs Phone: (973) 579-7227 www.thorlabs.com/ |  | 
Centration (arcmin) | <3 |  | 
Clear Aperture (% of diameter) | >90 |  | 
Design Diameter (mm) | 25.4 |  | 
Design Wavelength (nm) | 486.1 | 587.6 | 656.3
Diameter Tolerance (mm) | +0.0/-0.1 |  | 
Focal Length (mm) | 30 |  | 
Surface Quality | 40-20 Scratch-Dig |  | 
Thickness Tolerance (mm) | ±0.15 |  | 
Weight (kg) | 0.04 |  | 
Cost ($) | 77.00 |  | 
Ground Shipping – 1 week ($) | 8.75 |  | 

## Slide 7: Lens Assembly
Product Used in Design | C-Mount Lens Holder
Manufacturer | Thor Labs Phone: (973) 579-7227 www.thorlabs.com/
Price | $27.60
Shipping | $ 8.75 (6-9 days)
Size | Diameter = 42.164 mm Height = 9.525 mm
Quantity Required | 1
Simply insert lens into c-mount using twisting action. C-mount fits in microscope.

## Slide 8: Design: Electrical System _ PIN Diode Selection
Responsivity at 600 nm wavelength
Low Noise: Thermal Noise + Dark Current Noise
Low Capacitance
Operating Temperature
Size of active area
Notes: Explain the Reasoning.

## Slide 9: Design: Electrical System _ PIN Diode Selection
PIN-HR(S)008(L) Specification
Product Used in Design | PIN-HR(S)008(L)
Manufacturer | UDT Sensor. Inc Phone: (310) 978-0516 http://www.udt.com
Price | $22.00 to $246.00** (Price varies by the additional options)
Shipping | $ 9.28 (6-9 days)
Size | Diameter = 0.4572 ± 0.051 mm Height = 1.5875 mm (Body + Wire) = 0.3175 mm (Body)
Quantity Required | 1

## Slide 10: Design: Electrical System _ Input Constraint
Why?
  In Real OP Amps, small current (μA ~ pA) flows inside of OP Amps, generating Voltage error.
Solution: Current to Voltage Converter
Notes: Nah~~ phew pA error, who care? -> We do care

## Slide 11: AC Current to Voltage Converter
Literature Search: A low-noise and wide-band ac boosting current-to-voltage ampliﬁer for scanning tunneling microscopy
  Tunneling Microscopy
  Specialized current to voltage converter for pA to nA Current.
  Unity Gain at 5 kHz to 10 kHz

## Slide 12: AC Current to Voltage Converter
Product Used in Design | OPA111BM | OPA2111KP | Tefalon Standoff – PTFE Insulated Terminal Pins #11067
Manufacturer | Burr Brown Corp. Phone: 520) 746-1111 http://www.burr-brown.com | Burr Brown Corp. Phone: 520) 746-1111 http://www.burr-brown.com | Keystone Electronics. Corp http://www.keyeleco.com
Price | $53.35 | $15.23 | $0.94
Shipping | $2.95 (4-6 days) | $2.95 (4-6 days) | $4.99 (3-7 days)
Size | 9.08 x 9.08 x 4.4 mm3 | 9.3 x 6.5 x 6.6 mm3 | 4.37 Diameter x 6.35 mm3
Quantity | 1 | 2 | 1
Picture of the Product |  |  | 

## Slide 13: Design: Electrical System _ Filter
Domain of Frequency of interest: 5 kHz to 10 kHz
Preserve the shape of the original signal!
Solutions
  Analog Filter
    Active Bandpass Filter: Bessel Filter
  Digital Filter

## Slide 14: 4th order Bessel Bandpass Filter Design
Center Frequency FM= 7.5KHz
Bandwidth B= 5kHz
Q= FM/B = 1.5
Center Gain Km = 1 (absolute value); it’s is an unity gain filter
From the Coefficient of the 4th order Filter Table a1 = 1.3617 b1 = 0.6180 α = 1.2711 (at Q = 1.5)
Fm1 = FM/ α = 5900.4 Fm2 = FM* α = 9533.3
C = 10 nF

## Slide 15: 4th order Bessel Bandpass Filter Simulation
Design Failed

## Slide 16: Filter – Digital Filter
Axon™ Axopatch™ 200B microelectrode Amplifier
  Bessel & 4-pole Butterworth Filter
  Bandwidth: 30Hz to 5kHz
  Delay: Unknown

## Slide 17: OP Amplifier Design
Inverting Amplifier
Product Used in Design | OPA2111KP
Manufacturer | Burr Brown Corp. Phone: 520) 746-1111 http://www.burr-brown.com
Price | $15.23
Shipping | $2.95 (4-6 days)
Size | 9.3 x 6.5 x 6.6 mm3
Quantity | 1
Picture of the Product | 

## Slide 18: Assembly of the Circuits
Soldering on Printed Circuit Borad (PCB)
Product Used in Design | Multi-purpose OP AMP PCB
Manufacturer | AMZFX® http://www.muzique.com
Price | $10.00
Shipping | $ 2.95 (4-6 days)
Size | 31.75 x 48.26 x 4.76 mm3
Quantity Required | 4
Picture of the product | 

## Slide 19: Assembly of the Circuits
Schematics for Multi-purpose OP AMP PCB
PCB will be drilled and screwed to the cylinder in which adopted into the SM1 Adopter of the head-stage.

## Slide 20: Peltier Cooler Analysis
Normally calculation is performed based on an expected thermal load in Watts.
In our case, there is no load. The PIN diode must be cooled relative to ambient temperature.
Team discussed this with an engineer at Tellurex, a Peltier Cooler Company.
His advice was to pick a Peltier that fit our size constraints, and to then adjust the cooling amount by increasing or decreasing the DC current applied to the cooler.

## Slide 21: Design: Peltier Cooler Selection
 | Weight | C2-04-0101 | C2-06-0902R | 00801-9X30-10RU3 | C2-06-0402R | C2-04-0102 | 00301-9X30-10RU2
Qmax | 7 | 5 | 10 | 4 | 8 | 5.5 | 3
dTmax | 6 | 8 | 8 | 8.5 | 8 | 8 | 8.5
A | 10 | 7 | 6.5 | 9 | 7 | 8.5 | 10
B | 10 | 7 | 3 | 9 | 7 | 8.5 | 10
H | 10 | 7 | 7 | 9 | 7 | 8.5 | 10
 | Total | 230 | 283 | 349 | 314 | 341.5 | 372
Notes: Most important was dimensions. Qmax and dTmax are parameters of the cooler, on how much it can maximally cool.
A,B, and H are the dimensions, and are most important given our size constraint.

## Slide 22: Cooler Specifics
Product Used in Design | 00301-9X30-10RU2 Single Stage Mini Peltier Cooler
Manufacturer | Custom Thermoelectric, Inc. Phone: (443)-926-9135 http://www.customthermoelectric.com/index.htm
Price | $31.50 (includes moisture seal and lead wires)
Quantity | 1
Picture of the Product | 
Shipping Cost | $8.73
Shipping Time | ~5-8 days

## Slide 23: Mounting/Heat Sink/Thermal Interface

## Slide 24: Design Safe
Lack of Grounding.
Improper Wiring
Inadequate Heating/Cooling of Peltier (Heat Sink)
Wastes: Defective Parts

## Slide 25: Conclusion
Failed to properly design Bessel Filter.
All other design requirements were properly attained.
No IP.

## Slide 26: Team Responsibilities
Team Member | Specific Role
Eric Kleinberg | Research and Design of Peltier (TEC) Cooling System Website Development Schematic Development
Leran Firer | Research and Design of Optical Focusing System Lens Selection Designsafe Analysis
Dohyun Kim | Research and Design of Electrical Circuitry PIN Diode Selection Filter Simulation.

## Slide 27: Questions?